| Espiner — Bag Strength Test Development Bag Clamp Fixture ¦ Design Session Progress Report | PREPARED BY Suhas Sundaresh K — simple design people PERIOD ¦ MACHINE 16–24 March 2026 ¦ Mark-10 F755 EM |
| --- | --- |

CONFIDENTIAL · ISO 13485:2016 · All materials SS 316 throughout · Full material certs required

A full design session was completed for the mechanical clamping fixture holding a ripstop nylon conical bag during destructive compression testing at 590 N. All component geometries are defined, key design decisions are locked, and structural calculations are verified. The design is ready for Fusion 360 CAD finalisation — one primary measurement task remains before machining release.

| ACCOMPLISHMENTS THIS SESSION Bag geometry confirmed — Clamp plane Ø67.8 mm, upper ref Ø73.2 mm, taper ~8.1° half-angle Clamping principle finalised — Conical SS surfaces + 1 mm Shore 20A silicone liners. Dual grip: threaded rod (torque-controlled) + self-energising wedge at 8.1° self-locking boundary All 7 component dimensions defined — Bottom plate, ×4 standoffs, top plate, split collar, support ring, silicone liners, closing rods Structural calculations verified — All safety factors confirmed (see table) 5 documents produced — ENG-CLJIG-001 Rev E, weekly report, CAD dimensions ref, ring cross-sections, this summary 9 hours effort — 5 hrs CAD + 4 hrs design development STRUCTURAL CALCULATIONS VERIFIED CHECK; RESULT; SF. Standoff stiffness; k = 2,897 N/mm → δ = 0.19 mm at 560 N; —; ✓. Standoff fatigue; 14.1 MPa bending vs 120 MPa endurance; 8.5×; ✓. Ring wall (min); 3.4 mm wall, thick-walled cylinder; 9.5×; ✓. Friction grip; 697 N grip > 590 N test load; >1×; ✓. Dovetail root stub; 16 mm — adequate; —; ✓. Assembly clearance; 2.5 mm each side — elastomer through bore; —; ✓. | KEY DESIGN DECISIONS LOCKED Dovetail Retention Through male dovetail rail (60°, 12 mm). Collar slides into plate — shoulders resist uplift at bag failure. Zero retention fasteners.; Leadscrew Closing Half A = fixed jaw. Half B driven by ×2 M8 rods. Guided by dovetail walls — cannot skew regardless of tightening sequence.. Half A Adjustability M6×0.75 fine-pitch screw + slotted lock bolts. Bore centreline alignable with machine axis. Position recorded in DHF per session.; Bore Surface Finish Ra 3.2–6.3 µm — deliberate roughness for silicone RTV mechanical key. As-machined, no finishing pass — reduces cost.. Base Plate Fixing 5-bolt pattern: 4× #10-32 UNF + 1× ½-20 UNC. Countersinks on top face. No Mark-10 modification required.; Elastomer Cone — Rejected Cone on TP to stabilise ring rejected. Ring already stabilised by clamped bag wall. Revisit after first test data only.. PENDING BEFORE MACHINING RELEASE ITEM; ACTION REQUIRED; PRIORITY. Bag taper angle; Measure 5+ bags at 4+ heights — all dims depend on this; BLOCKER. Top plate dims; Confirm rectangle dimensions in Fusion 360; CAD. Half A fixing holes; Confirm positions — clear of dovetail & rod holes; CAD. M8 rod hole positions; Min 12 mm clearance above/below bore edges; CAD. Standoff bolt circle; Maximise spread within plate footprint; CAD. Clamping torque; Establish from first test — suggested 5 Nm per rod; Test. Rod & fastener lengths; Calculate from assembled CAD model; CAD. |
| --- | --- |

| IMMEDIATE NEXT STEP | Measure 5+ bag samples at 4+ heights each to confirm taper angle. This single task unblocks all remaining bore and ring dimensions and enables full machining release. Current 8.1° is derived from 2 measurements on 1 bag only — insufficient for production. |
| --- | --- |

| Espiner — Bag Strength Test Development ¦ Bag Clamp Fixture ¦ 16–24 March 2026 | ENG-CLJIG-001 Rev E ¦ simple design people ¦ CONFIDENTIAL |
| --- | --- |